Megohmmeter-Ohmmeter RL 2200

The RL2200 is particularly suited to maintenance, verification and reception of telephone lines, and more generally of all low voltage cables.

Its unique dual display offers both digital readings, for high accuracy measurements, and analogue readings to «feel» measurements tendency and monitor threshold limits.

The RL2200 measures loop resistance, insulation resistance, AC and DC voltages. Threshold limits can be set on resistance and insulation measurements.

- Dual analogue/digital display

- Multi-purpose

- Quick cable loading feature

- Auto-discharge feature

- Rugged and designed for field use

Functions

Insulation resistance measurements

Ranges and measurement voltages:

- from 0.01 MΩ to 1 000 MΩ under 50 V

- from 0.1 MΩ to 10 000 MΩ under 500 V (500 V constant voltage for insulation > 5 MΩ).

- measuring current: < 1 mA.

Insulation threshold indication (beep) above adjustable limit.

France Télécom approved

Resistance range under 50 V.................................................................................

| Analogue range | Digital range |  |  | Accuracy |  |
| --- | --- | --- | --- | --- | --- |
|  | Range | Measured resistance | Resolution | Analogue | Digital (1) |
| 0.03 MW to 3 000 MW | 0.00 MΩ to 99.50 MΩ | 0.00 to 10.00 MΩ 10.00 to 30.00 MΩ | 0.01 MΩ 0.05 MΩ |  | 5% + 0.01 MΩ |
|  |  | 30.00 to 50.00 MΩ 50.00 to 99.50 MΩ | 0.10 MΩ 0.50 MΩ | ± 1 segment |  |
|  | 100.0 MΩ to 1 000.0 MΩ | 100.0 to 200.0 MΩ 200.0 to 400.0 MΩ 400.0 to 700.0 MΩ 700.0 to 1 000.0 MΩ | 1.0 MΩ 5.0 MΩ 10.0 MΩ 20.0 MΩ |  | 10% |

(1) Accuracy is given 1 year in ±(% of reading + nΩ) at 23±1°C

Resistance range under 500 V................................................................................

| Analogue range | Digital range |  |  |  | Accuracy |
| --- | --- | --- | --- | --- | --- |
|  | Range | Measured resistance | Resolution | Analogue | Digital(1) |
| 0.3MΩ to 30 000 MΩ | 0.0 MΩ to 995.0 MΩ | 0.0 to 100.0 MΩ 100.0 to 300.0 MΩ | 0.1 MΩ 0.5 MΩ |  | 5% + 0.1 MΩ |
|  |  | 300.0 to 500.0 MΩ 500.0 to 995.0 MΩ | 1.0 MΩ 5.0 MΩ | ± 1 segment | 10% |
|  | 1000 MΩ to 10 000 MΩ | 1 000 to 2 000 MΩ 2 000 to 4 000 MΩ | 10 MΩ 50 MΩ |  |  |
|  |  | 4 000 to 7 000 MΩ 7 000 to 10 000 MΩ | 100 MΩ 200 MΩ |  |  |

(1) Over 1 year in ±(% of reading +nΩ) at 23±1°C

Resistance measurements

- Range : from 0 to 10 000 Ω

- Measuring current: < 1mA.

Continuity threshold indication (beep) below adjustable limit.

| Analogue range |  |  | Digital range |  | Accuracy (1) |
| --- | --- | --- | --- | --- | --- |
|  | Range | Measured resistance | Resolution | Analogue | Digital |
| 0.3 Ω to 30 000 Ω | 0.0 Ω to 999.5 Ω | 0.0 to 100.0 Ω 100.0 to 500.0 Ω 500.0 to 999.5 Ω | 0.1 Ω 0.2 Ω 0.5 Ω | ± 1 segment | 0.5% + 0.2 Ω |
|  | 1 000 Ω to 10 000 Ω | 1 000 to 2 000 Ω 2 000 to 3 000 Ω 3 000 to 5 000 Ω 5 000 to 10 000 Ω | 1 Ω 2 Ω 5 Ω 10 Ω |  |  |

(1) At 23±1°C over one year and in ±(% of reading + nΩ) on digital display.

DC voltage measurements

3 digits digital display.

Range 0 to 100 V DC, resolution 0.1 V, accuracy ±(1%rdgs. + 0.5 V).

Range 100 to 500 V DC, resolution 1 V, accuracy ±(1%rdgs. + 1 V).

Input impedance: 200 KΩ.

AC voltage measurements

Range 0 to 100 V AC, resolution 0.1 V, accuracy ±(1%rdgs. + 0.5 V).

Range 100 to 400 V AC, resolution 1 V, accuracy ±(1%rdgs. + 1 V).

Input impedance: 200 KΩ.

Average value translated to RMS without DC component.

Band width: 40 to 400 Hz (higher frequencies can be measured with lower accuracy).

General specifications

Dual display ........................................

- analogue display: 67 segments bargraph on 5 decades logarithmic scale.

- digital display: 4 digits LCD.

Adjustment of ¥ value to adapt to peculiar measurement situation.

Safety.................................................

Using insulation measurement function under 50 V and 500 V, measuring voltage is applied on line only while measuring key is pressed. When the key is released, the cable is automatically unloaded within 1sec. Overload protection up to 380 V RMS on all ranges (90 V RMS on resistance ranges).

Quick load ...........................................

Measurement settling time < 2 sec. for line capacitance < 1μF.

Threshold limits ...................................

Set up of limits on resistance and insulation ranges with audible warning (beep).

Presentation........................................

Plastic housing.

Dimensions : 195 x 100 x 45 mm.

Mass: 0.500kg.

Protection: IP54.

Connection: 2 safety 4mm plugs.

Power supply ......................................

4 x R6 or LR6 1.5 V type batteries.

Autonomy: 4500 measurements of 5 sec.

each under 500 V with resistance ³ 0.5 MΩ.

Auto switch-off function (after 30mn).

Environment ........................................

Operating range limits: -10 to +50°C, 20 to 80% relative humidity without condensation.

Storage and shipping range from -30°C to + 50°C.

Included as standard ............................

- Carrying case for «hands free» use

- Measuring leads and «crocodile» type clamps.

Ordering instructions

Analogue/digital Megohmmeter-ohmmeter RL 2200