Graphic stand alone, self powered datalogger SA 32

The SA32 is intended to measure, handle, monitor and log record analogue signals. Designed with both flexibility and precision in mind, this system can be used in numerous applications such as:

- data collection in the field

- research or quality laboratories

- acquisition interface for calculator or computer

- process monitoring.

Different versions of the SA32 are available:

• With keypad and display

• With keypad, display and PCMCIA interface,

• Without keypad, display neither PCMCIA interface.

Universal programmable inputs

600,000 counts display accuracy 1μV resolution

Enhanced calculation & statistics functions

Bulk storage memory

PC software included

Functions

SA32 allows programming of 100 addresses, which can be defined to measure real channels (up to the system limit of 32 channels max) or virtual channels used to perform calculations/maths functions.

Each channel can be assigned a 4 character engineering units label, an 8 character mnemonic and the following functions:

• Monitoring

4 alarm limits per channel, different for each channel. SA 32 is equipped with 2 output relays addressable to any limit.

• Maths functions

Available as standard are the operators +, -,/,x, elevation to multiple powers, timer function(s), logic operators, square root, log decimal, absolute value and statistical functions min, max, average, standard deviation.

• Conditional functions

Such as IF, THEN, DO allowing conditional modification of parameters and change of scanning speed or of criteria for storage in memory.

Scanning .............................................

Input channels are scanned by a CMOS logic circuit.

Scanning is organised in «tasks» or scanning procedures, defining what to scan (channels), how to scan (scanning start and stop, scanning rate), and how to process the results (mass storage, RS232 output). The user has access to 16 tasks, therefore 16 different scanning subgroups.

Programming.......................

The logger is programmable entirely via the keyboard and help menu on screen, or via RS232 interface with the assistance of LS32 DOS software supplied with the unit.

Display ...............................................

Backlit LCD give simultaneous display of 2 channels in either digital or graphical format (150 points definition) on a time base adjustable between 1 sec. to 1 min. It also permits recall of the storage memory in order to display recorded data in digital or graphic form.

Configuration

Inputs Input number

From 8 to 32 inputs are possible, depending on the type of connection (1, 2, 3 or 4 wires). A supplementary 4 wire input is available on the front panel.

These inputs are subdivided in 8 subgroups of 5 wire. Each of the sub-groups has a common (isolated from the other sub-groups) and can be programmed independently to perform different types of measurements:

4 x 1 wire (single wire)

2 x 2 wires (differential)

2 x 1 wire and 1 x 2 wire (2 single wire + 1 differential)

1 x 4 wires (4 wire measurement).

Note: for 3 wire measurement, it is necessary to make a strap on the connector.

Max admissible voltage between 2 channels is 50 V.

One pulse counting input is available, allowing to summarise pulses delivered by a dry contact.

Max. frequency: 50 Hz.

Minimum closing and opening duration: 10msec.

Sampling rates

Three sampling rates are programmable channel by channel. Sampling rate affects settling time and conversion rate, as the highest rate gives the least resolution and precision (multi-ramp converter).

| Sampling rate | Integration time | Display capacity | Number of counts | UR (1) |
| --- | --- | --- | --- | --- |
| 7 m/s (2) 20 m/s 50 m/s | 100 ms 20 ms 1 ms | ± 780 000 p ± 78 000 p ± 7 800 p | 390 000 78 000 7 800 | 2 1 1 |

(1) UR = units of representation (as per IEC 485: minimum gap between 2 successive values)

(2) At 7 meas/sec, 60 mV and 60 W ranges are measured at 78 000 counts and 1 UR. Resistance and Pt 100 W ranges in 3

wires are measured in 2 times. Sampling rates are half of those indicated above.

Types of measurement

DC voltage ...............................................................................................................................

| Range (1) | Resolution |  |  | Accuracy at 7 m/s (2) |  | Accuracy at 20 m/s |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | 7 m/s | 20 m/s | 50 m/s | 90 days | 1 year | 90 days | 1 year |
| 60 mV 600 mV 6 V 0 - 10 V | 1 μV 2 μV 20 μV 20 μV | 1 μV 10 μV 100 μV 100 μV | 20 μV 100 μV 1 mV 1 mV | 0.01% + 4μV 0.01% + 8μV 0.01% + 80μV 0.01% + 80μV | 0.02% + 6μV 0.02% + 10μV 0.02% + 0.1mV 0.02% + 0.1mV | 0.01% + 4μV 0.01% + 20μV 0.01% + 0.2mV 0.01% + 0.2mV | 0.02% + 6μV 0.02% + 30μV 0.02% + 0.3mV 0.02% + 0.3mV |

(1) 30% of scale over-range permissible.

(2) Accuracy is given as ± (% of reading + n units) at 23 ±1°C.

Reproducibility between 2 different channels is < 2 C (C being the accuracy expression constant on 90 days). For example for 60 mV range at 7 meas/sec, C = 4 μV.

Input impedance >1000 MΩ.

Input current < 5 nA (typical: 300 pA)

Serial mode rejection rate (60 mV range)

> 70 dB at 50 and 60 Hz at 7 meas/sec

> 70 dB at 50 Hz at 20 meas/sec.

Maximum common mode admissible voltage: 50 VDC or 50 VAC peak.

Common mode rejection rate (60 mV range)

> 120 dB DC and AC at 50 and 60 Hz at 7meas/sec

> 120 dB DC and AC at 50 Hz and at 20meas/sec.

Maximum admissible voltage between different channels: 50V DC or 50 V AC peak.

Temperature coefficient < 10% of accuracy/°C.

Resistance........................................................

| Range (1) | Resolution |  |  | Accuracy at 7 m/s (2) |  | Accuracy at 20 m/s (2) |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | 7 m/s | 20 m/s | 50 m/s | 90 days | 1 year | 90 days | 1 year |
| 60 W 600 W 3 000 W | 1mΩ 2mΩ 20mΩ | 1mΩ 10mΩ 100mΩ | 10mΩ 100mΩ 1Ω | 0.02% + 4mΩ 0.02% + 8mΩ 0.02% + 80mΩ | 0.025% + 6mΩ 0.025% + 10mΩ 0.025% + 0,1Ω | 0.02% + 4mΩ 0.03% + 20mΩ 0.03% + 0,2Ω | 0.025% + 6mΩ 0.025% + 30mΩ 0.025% + 0,3Ω |

(1) 30% of scale over-range permissible.

(2) Accuracy is given as ± (% of reading + n units)at 23 ± 1°C in 4-wire mode.

Accuracy is given in 3 wires/balanced configuration.

In case of wire unbalance, accuracies must be increased by the accuracy expression constant on 90 days plus line unbalance value.

Measuring current: 1mA

Allowable line resistance 100W per wire.

Temperature by Pt 100W.................................................................................

| Scale | Resolution |  |  | Accuracy at 7 m/s (1) |  | Accuracy at 20 m/s (1) |  | Accuracy at 50 m/s (1) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 7 m/s | 20 m/s | 50 m/s | 90 days | 1 year | 90 days | 1 year | 90 days | 1 year |
| - 220 to 850°C | 0.01°C | 0.1°C | 1°C | 0.02% + 0.1 | 0.025% + 0.1 | 0.02% + 0.12 | 0.025% + 0.15 | 0.03% + 0.6 | 0.04% + 0.6 |

(1) Accuracy is given as ±(% of reading + n °C) at 23°C ±1°C in 4 wire mode.

Accuracy is given for 3 wires/balanced configuration. Accuracies must be increased by accuracy expression constant on 90 days.

The sensor uncertainty must be taken into account.

Thermometric resistance 100 W at 0°C as per IEC 761 (DIN 43760).

Measuring current: 1mA.

Allowable line resistance 100 W per wire.

Reproducibility between 2 different chan - nels < 2 C.

Temperature coefficient < 0.003%/°C.

Temperature with thermocouple ...............................................................................................................................

Cold junction compensation is performed by a platinum probe. When used with an internal reference junction, the errors should be increased by 0.5°C. With an external reference junction the errors should be increased by 0.1°C and by the error of the probe used.

Reproducibility between 2 different channels: < 2 C (C being the accuracy expression constant over 90 days. For example, for type K thermocouple at 7 meas/sec, temperature below 100°C, C = 0.2°C).

Input impedance: > 1000 MW

Input current: < 5 nA (typical 300 pA).

Maximum admissible line resistance: 100W per wire.

Temperature coefficient < 10% of accuracy/°C.

Contact ...............................................

Dry contact (potential free).

Measuring current 1mA.

Level 0: line resistance Ro > 100 kW

Level 1: line resistance R1 < 1 kW.

Process current ranges 0-20 mA and 4-20 mA .............................................

These measurements are performed using external 50 W shunts.

| Ty p e | Resolution |  |  | Range | Accuracy at 7 and 20 m/s (1) |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | 7 m/s | 20 m/s | 50 m/s |  | 90 days | 1 year |
| K | 0.5°C 0.2°C 0.1°C | 0.5°C 0.2°C 0.1°C | 1°C 1°C 1°C | - 250°C to - 200 °C - 200°C to - 100 °C - 100°C to +1370°C | 0.01% + 1 °C 0.01% + 0.4 °C 0.01% + 0.2 °C | 0.02% + 1.4°C 0.02% + 0.5°C 0.02% + 0.3°C |
| T | 0.5°C 0.2°C 0.1°C | 0.5°C 0.2°C 0.1°C | 1°C 1°C 1°C | - 250°C to - 200 °C - 200°C to - 100 °C - 100°C to + 400 °C | 0.01% + 0.8 °C 0.01% + 0.4 °C 0.01% + 0.2 °C | 0.02% + 1 °C 0.02% + 0.5°C 0.02% + 0.3°C |
| J | 0.2°C 0.1°C | 0.2°C 0.1°C | 1°C 1°C | - 210°C to - 120 °C - 120°C to +1100°C | 0.01% + 0.3 °C 0.01% + 0.2 °C | 0.02% + 0.4°C 0.02% + 0.3°C |
| S | 1 °C 1 °C | 1 °C 1 °C | 1 °C 1 °C | - 50°C to + 120°C + 120°C to +1768°C | 0.01% + 1.2 °C 0.01% + 0.7 °C | 0.02% + 1.7°C 0.02% + 1 °C |
| B | 1 °C 1 °C | 1 °C 1 °C | 1 °C 1 °C | + 400°C to + 900 °C + 900°C to +1820°C | 0.01% + 1.5 °C 0.01% + 0.8 °C | 0.02% + 2 °C 0.02% + 1 °C |
| N | 1 °C 0.2°C 0.1°C | 1 °C 0.2°C 0.1°C | 1°C 1°C 1°C | - 250°C to - 200°C - 200°C to - 120°C - 120°C to +1300°C | 0.01% + 1.5 °C 0.01% + 0.6 °C 0.01% + 0.3 °C | 0.02% + 2 °C 0.02% + 0.9°C 0.02% + 0.4°C |

(1) Accuracy is given as ± (% of reading + n°C) at 23 ±1°C when using a reference junction at 0°C.

Outputs

Dry relay outputs ................................

SA32 is equipped with 2 output relays each of which has a common and one working contact on the rear terminal

block.

Interrupt capability with resistive load: 48 V or 1A or 30 W.

Life expectancy 5 x 105 operations.

These relays are operated from various stimuli, such as:

- specific alarm threshold,

- conditional processing,

- or via RS232 interface.

Analogue output ..................................

SA32 has an analogue output 0 to 2.55 V, galvanically isolated from the inputs.

Analogue output resolution is 256 counts.

Accuracy is ± 10 mV for a resolution of 10mV.

Minimum admissible load resistance 2.5 kW.

This analogue output allows retransmission of any of the 100 addressable channels.

Communications interface

The communication interface is programmable either standard RS 232C or with MODBUS RTU protocol.

It is accessed on a 9-pin female SUB D connector.

The logger can handle directly one printer, one VT100 compatible VDU, one calculator programmed in MODBUS RTU protocol, it can also communicate directly with a PLC or supervisory PC.

All programming instructions can be sent to the logger via the interface.

Speed: 9600 bauds in standard RS 232C or protocol MODBUS RTU.

Character format: 8 bit, no parity, 1 stop bit.

MODBUS RTU is a registered trademark of AEG Modicon Inc

General specification

Display ...............................................

Retro-lit graphic screen (240 x 64 pixels), display of 2 channels simultaneously in numeric or graphic form.

Operating Conditions ...........................

Reference temperature: 23 ± 1°C.

Relative humidity (RH): 45 - 75%.

Nominal ambient range: 0 - 50°C, 20 - 75% RH.

Supply ................................................

Mains 110 to 240 V ±10%, 50 to 400 Hz.

Rechargeable battery incorporated.

Presentation........................................

Dimensions: 225 x 88 x 310mm; weight 3 kg.

ABS case for bench use.

Rack mounting option.

Software

An utilities library for PC compatible computers is available under Windows® 3.11 to NT in English and French. These utilities simplify the use of the equipment as well as widening its scope.

Operation and storage program LS32W

This routine enables the logger to be programmed from a computer; it is presented in the form of a scrolling menu with windows with the operator being guided through the options by a question and answer system.

LS32W also authorises the transfer of information stored in the logger to a PC.

The file created can then be converted with the utilities to a format compatible with spreadsheets (e.g. Lotus, Excel ...) or to a graphical representation of trends.

Communication between logger and PC can be performed by modem.

LS 32W programming screen

The Menu guides the operator right through the defining options e.g.

- sensor type

- conversion formulae

- alarm limits.

Graphics screen

LS 32W allows transfer of information stored in the memory of the logger into the computer; a back-up file is created and can be recalled for later work.

In the standard version, the user has access to graphic representation in the form of a curve vs time or as correlation between channels.

It is possible to define up to 30 simultaneous channels (these are distinguished by different colours with associated

mnemonics) and 2 separate scales.

Results produced can be printed on compatible printers (Epson or IBM, HP Deskjet 500, 560C).

Real time display and telemonitoring ....

Many Dos and Windows® software allow real time data display as graphs, bargraphs and numeric tables. It is also possible to create datafiles and to use them for later processing.

File structure .......................................

Files created by LS32W software are of binary type - various sort options are offered (sub-files extracting certain channels, by marker position, by choosing channels in alarm state etc...)

The sub file(s) can be directly converted before being imported into the usual spreadsheets Lotus, Excel, D-Base, Symphony etc.)

Version

PC memory card version ......................

This version allows data storage as well as configuration saving and loading.

SA32 data loggers with keypad and display can be equipped wit a PCMCIA interface allowing data storage on static RAM or ATA memory cards (PCMCIA type 2).

These cards allow the configuration and result files to be loaded or saved.

Being PC compatible these files can be processed on every computer equipped with a PCMCIA interface and LS32 or LTC32W software. In the same way, every configuration files created on a computer can be saved on the card and downloaded in the logger.

The number of files created by the user is not limited, but the data logger can only handle 16 result files simultaneously.

Note: if the PC has not got a PCMCIA interface, the card can be read either via SA32 or via an external PCMCIA drive connected to the parallel interface.

One measurement takes 6 bytes.

A 10 Mbytes card can store 1 650 000 measurements.

Ordering instructions

Datalogger with keypad, display and with 14 000 data memory               SA 32 C09

Datalogger with keypad, display and 64 000 data memory                   SA 32 C58

Datalogger without keypad and display and with 14 000 data memory          SA 32 A09

Datalogger without keypad and display and with 64 000 data memory          SA 32 A58

Datalogger with PCMCIA interface SA32XXXP

Accessories ....................................................................

PC RS485 board                                                    AN5891

PC RS232/RS485 interface                                            ATC022

Transport case                                                      AN 6901

Clipsable connector                                                  ER 48358-000

Connector 4 pins                                                    ER 48345-004

Shunt                                                             ER 44007-024

Panel mounting kit                                                  AN 5883

Rack mounting kit                                                   AN 5884

RS232 (9-25 pins) for printer                                          AN 5876

Serial printer (40 columns)                                            ATC 001

Set of ten paper rolls for ATC001                                      ATC 002

Attenuator                                                        ATC 005

PCMCIA drive                                                     ATC012

ATA PCMCIA card 2 Mb                                            ATC013

ATA PCMCIA card 7 Mb                                            ATC014

ATA PCMCIA card 15 Mb                                           ATC015

Static RAM PCMCIA card 1 Mb                                       ATC016

Windows programming software                                       LS32W

Windows software                                                  LW1

DLL library                                                       LTC001

Labview driver                                                     LTC003

Other software                                                     On request

Calibration certificate or other documentation                            On request