DIGITAL PANEL METERS ICO 61

Electrical measurements 10 000 points

I CO 61 are digital panel meters of which the function of measurement,range, thresholds and of the specific scalings are configurable.

The configured parameters are protected from the inopportune modifications.

- Format 96 x 24 mm

- Function, ranges, programmable thresholds and scale.

- 2 thresholds with exit relay

Functions

Measure...............................................

The ICO61 allow measurements of A.C/D.C voltage 400V (true effective value with or without component continuous health) and measurements of A.C/D.C current up to 5A and measurements of frequencies up to 5000 Hz.

They comprise:

- a power supply for standard sensor Namur,

- the blocking of measurement,

- the memorizing of the maximum and the minimum of a series of measurements.

Automatic change of range...............................................

In the same measuring unit.

Scaling...............................................

Whatever the function, the user can program a scaling which is clean for him.

The value to be displayed is entered using the keyboard, while the value of corresponding measurement can either be introduced in the same way or measured by the instrument.

Thresholds...............................................

Programming of two thresholds and their hysteresis.

Capacity breaking of the associated relay 250 V - 1 A or 60 VA to 50 Hz, 150 V - 1 A or 30 W in continuous.

Operation is in agreement with the safety requirements ,alarm being characterized by the not fed repercussion of the relay correspondent (coil).

(1) Extended from measurement : of - 40 to + 110 % range

(2) ± (% reading+ a number of units)on 1 year Continuous : de - 40 à 100 % range Into A.C : of 4 to 100 % range from 25 to 500 Hz without continuous component (possible measurements up to 10 Khz on certain ranges with less accuracy)

(3) Measure of current with shunts

(4) Two levels for entry voltage : 50 mV~ à 50 V~ et 5 V~ à 250 V~.

In the course of measurement, each threshold can be visualized.

Circuits RC of protection are envisaged out of accessory for the protection of relay contact if those are connected to inductive circuits.

Protection of the programming...............................................

The instruments have two levels of protection by software.

l Protection of the unit of the programming by access code.

l Protection of the threshold values by programming of the function safety.

General characteristics

Display...............................................

From -1999 to 9999 with indication, LED red,yellow or green,of 14 mm.

Speed of measurement : 2,5 measures/s.

Temperature coefficient £ 10 % accur - acy/°C.

Acceptable maximum voltage on the rages voltage : 600 V~.

Frequency...............................................

Acceptable maximum voltage :- 70 V peak on low entry level (50 mV à 50 V), - 570 V peak on entry high level (5 V à 250 V).

Acceptable max. continuous components:

- 50 V- peak on low entry level (50 mV à 50 V),

- 400 V- peak on entry high level (5 V à 250 V).

Voltage of common mode maximum acceptable on the entries : 250 V~.

Common mode...............................................

Accepteble maximum voltage: 250 V ~.

Rejection of common mode ³ 110 dB to 50 Hz on the continuous gauge 500 mV.

Rejection of series mode ³ 60 dB to 50 Hz on the continuous gauge 500 mV.

Climatic Environnement...............................................

Nominal field : 0 to 50°C, of 20 to 80% HR without condensation.

Isolated power supply...............................................

230 V~ (± 10 %) 50 to 60 Hz.

Others power supply:

5 V- ± 10 %,

9 to 18 V -

18 to 36 V-,

36 to 72 V-,

24 V~ ± 10 %,

48 V~, ± 10 %,

115 V~± 10 %.

Presentation...............................................

Case ABS.

Sealing: IP 40.

Weight : approximately 300 g

Connection...............................................

Removable screwing connectors supplied with instrument.

Dimensions

Accessories

External shunts...............................................

Modules screwed on the connector

AN 8003 : 1 A, 100 mV, precision 0,2 %;

AN 8004 : 5 A, 100 mV, precision 0,5 %;

Dimensions : 25 x 27,5 x 14,8 mm.

Transformers of current...............................................

Classe 0,5.

AN 5844 : 2/1 A,

AN 5845 : 20/1 A,

AN 5846 : 200/1 A.

Sensors of standard proximity «NAMUR» (see corresponding documentation).

Protective circuit...............................................

Circuit RC of protection of the contacts alarm for cut on inductive circuits (250 V~ max.) : ER42062.

Instructions to order

Accessories and option...............................................

Shunt 1 AN 8003

Shunt 5 A AN 8004

Transformer of current 2/1 A AN 5844

Transformer of current 20/1 AN 5845

Transformer of current 200/1A AN 5846

Sensors of proximity IEA 2001to IIA 2015

Circuit RC ER42062