DIGITAL PANEL METERS TNE 2

Temperature measurements 2000 points

Digital panel meters TNE2 can be used with temperature sensors.

Temperature is displayed in °C.

Those industrial apparatus are automatically delivered with a removable screwing connector, and an analog output.

The large readable display is made of red LED (green or yellow in option)

- -200 to +1200°C

- Dimensions : 72 x 24 mm

- Linear analog output

- Display : 14 mm LED

Functions

Thermocouples Temperature, K or J……………………………….

(response curves according to NFC 42 321 (CEI 584/1)

(1) ± (% of reading + digit degrees) at 23°C +/- 1°C for one year

The correction of reference junction is into the screwing connector. Maximal uncertainty due to reference junction: 0.5°C; if external CJC: 0.3°C

RTD Temperature, Pt100……………………………….

(response curves according to NFC 43 760 (CEI 751)

(1) ± (% of reading + digit degrees) at 23°C +/- 1°C for one year

3 wires mounting.

Acceptable line resistance :

10 ohm per wire (maximum unbalanced between wires 0.04 Ω)

Measurement current in sensor < 0.4 mA.

General Characteristics

Display……………………………….

2000 points (3 ½ digits), 14 mm height red LED (green or yellow in option), setting of brightness.

The last 3 digits switch off when overload occurs

Automatic zeroing and polarity. Sampling rate……………………………….

2.5 measurement per second

Temperature coefficient……………………………….

≤ accuracy/10°C

Security in case of breakdown of sensor………………………………..

makes decrease the measurement down to minimum value

Linear analog output……………………………….

1mV/digit linearity +/- 0,1% acceptable load > = 2kohm non insulated analog output

Operating conditions……………………………….

Nominal conditions : 0 to + 50 °C

Limit conditions : -10 to + 50°C

Power supply……………………………….

230 V~ (± 10 %) 50 to 60 Hz

Other power supplies :

6 V- ± 20 % non-insulated,

5 V- ± 10 % insulated,

9 to 18 V- insulated,

18 to 36 V- insulated,

36 to 72 V- insulated,

24 V~ ± 10 % insulated,

48 V~ ± 10 % insulated,

115 V~ ± 10 % insulated.

Consumption……………………………….

4 VA max

Presentation……………………………….

ABS Case.

I IP 40

Weight : 300 g. approximately

Connections

Removable screw connector

Connection……………………………….

Removable screw connector

Options

Watertight metallic case

IP 65 Front panel watertight

Dimensions : 99 x 36 x 120 mm

Accessories

Mechanic adaptator……………………………….

Allows adaptation of VNE into slightly larger cutout Adaptation of 75x25mm panel Meters External dimensions 80 x 32 mm

Dimensions

Options and accessories……………………………….

Flange (72 x 24 mm) AMD001

Display meter 12 channels refer to manual BCM012