GeminiPLUS 550/700

Dry Block Calibrator

700ºC to 50ºC

The GeminiPLUS Dry Block range offers industry-leading performance in an easy to use portable package - ideal for the calibration of thermocouples and platinum resistance thermometers. It has been designed to have a large block suitable for more probes, or larger probes than can be accommodated in the Jupiter and Calisto models.

The GeminiPLUS is available with a fixed block with 8 pockets or with a large removable block, see Gemini LRI overleaf. The GeminiPLUS is available in two models, the BASIC (B) and the SITE (S). The B model includes a sophisticated temperature controller with a dual display for Set Temperature and Dry Block Temperature. The large block has the benefit of higher capacity but with longer heating and cooling times compared to the fast response models such as the Jupiter 650.

550ºC to 35ºC

The S model includes a built in digital indicator to which an external standard thermometer can be connected giving greater accuracy eliminating temperature gradient and loading errors. For lab accuracy the GeminiPLUS can be used with a high-end temperature indicator such as one of the Isotech True Temperature Indicators (TTI).

Includes as standard: Windows Software, Computer Interface and a Ramp to Set Point Feature.

Increased resolution of ±0.01 available throughout the range via the PC interface and from 0.01 to +99.99 locally on the auto-ranging front display. The controller features multi-point block to display correction giving good absolute accuracy.

The S model has universal sensor input allowing Platinum Resistance Thermometers, Thermocouples (Types K, N, R, S, L, B, PL2, T, J and E) along with Linear Process Inputs including 4-20mA current transmitters to be displayed on the inbuilt indicator. The indicator can be programmed with up to five calibration points to provide high accuracy digital probe matching.

The indicator and controller are both addressable over the communications link.

New! The Site model can now be used with the supplied Cal Notepad software to test thermostats.

Key Features

_ Calibrate Whole Calibration Loop - using a heat source rather than an electrical simulator a test instrument and probe can be calibrated as a system.

_ High Capacity Block - Large Calibration volume for many sensors and low loading errors.

_ Two Operating Ranges - Fast response from 35ºC to 550ºC or 50ºC to 700ºC.

_ Simple To Use - Outstanding Value for Money!

_ S Model Includes Universal Input – Temperature Indicator allowing for up to five “correction” points to be programmed.

_ Windows Software and PC Interface as Standard.

_ Free Evaluation Report - ask for full data or visit

www.isotech.co.uk/industrial

Options

| Metal Block Sleeves |  |  |
| --- | --- | --- |
| GeminiPLUS 550 | Please specify | Set of four Sleeves to suit the block. Optional single hole sizes 4, 6, 8, 10, 12, 14mm diameter all 150mm deep. |
|  | 857-07-01 | Undrilled sleeves for local machining |
|  | 857-07-03 | 1 sleeve with 2 holes 4.5mm x 150mm deep |
| GeminiPLUS 700 | Please specify | Set of four Sleeves to suit the block. Optional single hole sizes 4, 6, 8, 10, 12, 14mm diameter all 150mm deep |
|  | 857-07-02 | Undrilled sleeves for local machining |
|  | 857-07-04 | 1 sleeve with 2 holes 4.5mm x 150mm deep |
|  | Note: The use of sleeves will introduce an additional thermal gradient into the block. This can be avoided by using the LRI model with a block drilled for specific probes |  |
| UKAS Calibration | UKAS | Calibration available to Order |
| Standard Probe | 935-14-72 | Platinum Resistance Thermometer for use up to 650ºC |
|  | 935-14-63 | Type N Thermocouple for use up to 700ºC |
| Carrying Case | 931-22-64 | Sturdy case accommodates the unit with room for accessories |

Calibration and Uncertainty

A certificate, traceable to National Standards, is included as standard. Recommended is an optional UKAS five-point calibration.

The accuracy will depend very much on the mode of use and the types of sensor to be used. Please contact Isotech for tutorials and uncertainty calculations and comprehensive evaluation reports. The GeminiPLUS 550/700 meets the Calibration Capacity requirements of EA I0/13, “EA Guidelines on the Calibration of Temperature Block Calibrators”.

Temperature Range: GEMINIPLUS 550 35ºC to 550ºC

GEMINIPLUS 700 50ºC to 700ºC

Absolute stability over 30 minutes: At 35ºC ±0.05ºC

At 275ºC ±0.05ºC

At 550ºC ±0.05ºC

At 700ºC ±0.05ºC

Computer Interface: Included with Windows Software (page 74)

Cools from:    GeminiPLUS 550

550ºC to 275ºC:  35 minutes

550ºC to 60ºC:  345 minutes

Heats from: GeminiPLUS 550   GeminiPLUS 700

30ºC to 550ºC: 35 minutes       N/A

50ºC to 700ºC: N/A            110 minutes

Best Performance: See Graph

Calibration volume Standard Block: Four 8mm pockets, four 19.5mm pockets, all 160mm deep

Display Resolution: 0.01 to 99.99

0.1  100.0 to 700.0

PC can display 0.01 across whole range with the software included

Indicator units: ºC, ºF, K

Power: 100 to 120V (50 / 60 Hz) or

200 to 240V (50 / 60 Hz)

600 Watts

Overall dimensions: Height 302mm

Width 176mm

Depth 262mm

Weight: GeminiPLUS 550 8.5kg

GeminiPLUS 700 14kg

How to Order: GeminiPLUS 550 or GeminiPLUS 700

Please specify model type required

Please specify voltage required

Please specify options required